## AccessLevelExpert
Project | Web software for Spectro::lyser v3.0
Module Name
Reference Document | Product Requirement Document web software for spectro::lyser v3.0
Version
Test data | https://data.s-can.at/
Testcase ID | Functional Requirement -DataView | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW1TS001 | The main page of the user interface should provide numerical data visualization of the current parameter values | Must have | Login as expert password:bernd | 1.Enter the URL 2.Enter the valid login credentials 3.Click on Perform login 4.Navigated to Values page 5.Verify the current parameters values 6.Logout as Expert | Latest values should be displayed for all the parameters
SW1TS002 |  | Must have | Login as expert password:bernd | 1.Enter the url 2.Enter valid login credentials 3.Click on perform login 4.Navigated to the values page 5.Click on any parameter controllers from the values dashboard page 6.click on logout | After clicking on parameter controllers modal dialogue should have the value & parameter details.
SW1TS003 | Parameters indicating a problem (NaN) shall be highlighted. Further information on the problem source shall be provided in a human readable form | Must have | Login as expert password:bernd | 1.Enter the URL 2.Enter the valid login credentials 3.Click on Perform login 4.Navigated to Values dashboard page 5.Check for parameters indicating a problem 6.Logout as expert | Parameter should highlight a problem and qulality issue should be provide in a human readable form
SW1TS002.1 |  | Must have | Login as expert password:bernd | 1.Enter the URL 2.Enter the valid login credentials 3.Click on Perform login 4.Navigated to the values page 5.Click on any parameter controller 6.Check the Quality issue status in a value details pop up window 7.Logout | If parameter controlled is highlighted with red color & the qulaity issue should be displayed as "no medium detected" & "value below medium"
SW1TS002.2 |  | Must have | Login as expert | 1.Enter the URL 2.Enter the valid login credentials 3.Click on Perform login 4.Navigated to the values page 5.Click on any parameter controller 6.Check the Quality issue status in a value details pop up window 7.Logout | If parameter controlled is highlighted with blue color & the qulaity issue should be displayed as "everything OK"
SW1TS004 | Information on the timestamp of last measurement and the time until next measurement shall be shown | Must have | Login as expert password:bernd | 1.Enter the URL 2.Enter the valid login credentials 3.Click on Perform login 4.Navigated to Values dashboard page 5.Check the timestamp for last measurement 6.Logout as expert | Information about the time stamp of last measured value & the time until next measurement should be displayed
SW1TS005
SW1TS006
SW1TS007
SW1TS008
Testcase ID | Functional Requirement -SW2.Spectral Data | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW2TS001 | A View to show the spectral absorption shall be present | Must have | Login as expert password:bernd |  |  |  |  | Spectral absorption is not available in the link
Testcase ID | Functional Requirement -SW3.Time Series | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW3TS001 | A View to show historical parameter values shall be present.The content of the data logger should be displayed graphically as a time series | Must have | Login as expert password:bernd | 1.Enter the URL 2.Enter the valid login credentials 3.Click on Time Series 4.Navigated to time Series 5.Verify content of data logger by graph 6.LogOut | Graphical representaion of data logger content
SW3TS002 | Time Series view shall have pan and zoom options | Must have | Login as expert password:bernd | 1.Enter teh URL 2.Enter the valid login credentials 3.Click on Time Series 4.Navigated to time Series 5.Verify Pan options byclicking 6.LogOut | Time Series graph should pan and Zoom option
SW3TS004 | The User should have the option to filter the parameters that are dispalyed | Must have | Login as expert password:bernd |  |  |  |  | There is no option to filter the parameters under
SW3TS005 | The user should be able to see if parameter warnings are associated with a measurement value and see them in human readable form | Must have | Login as expert password:bernd
Testcase ID | Functional Requirement - Spectral Time Series | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW4TS001 | A view to show the historical spectral absorption in a 3D graph shall be present. | Must have | Login as expert password:bernd |  |  |  |  | Spectral absorption is not available in the link
3.2.Measurement Settings
Testcase ID | Functional Requirement -SW5.Parameter Management | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW5TS001 | The user shall have the possibility to rename parameters. | Must have | Login as expert password:bernd | 1.Enter the URL 2.Provide the Valid login credentials 2.Click On Service 3.Select Measurement Settings 4.Click on "Enter Service mode" 5.Click on edit symbol beside parameter name 6.Rename the parameter name in parameter properties box 7.Click on Ok button 8.Logout | User should be able to rename the parameters
SW5TS002 | The user shall have the possibility to change the unit of a parameter by entering a custom unit | Must have | Login as expert password:bernd | 1.Enter the URL 2.Provide the valid login credentials 2.Click On Service 3.Select Measurement Settings 4.Click on "Enter Service mode" 5.Click on edit symbol beside parameter name 6.Change the unit in parameter properties box 7.Click on Ok button 8.Logout
SW5TS003 | The user should have the possibility to change number of digits displayed for the parameter | Must have | Login as expert password:bernd |  |  |  |  | Not found in the measurement settings page
SW5TS004 | The user should have the option to filter and rearrange the order of the parameters that are displayed | Must have | Login as expert password:bernd |  |  |  |  | Not found in the link
SW5TS005 | The user should be able to reset or see the factory default values of changed settings | Must have | Login as expert password:bernd |  |  |  |  | Not found in the link
Testcase ID | Functional Requirement -SW6.Measurement speed and interval | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW6TS001 | The user gets the possibility to configure the interval for periodic measurements in the range between 10 seconds and 24 hours | Must have | Login as expert password:bernd | 1.Enter the URL 2.Login using Credentials 3.Click On Service 4.Select Measurement Settings 5.Click on "Enter Service mode" 6.Set range for measurement interval 7.Click on "Leave Service mode" 8.Logout |  |  |  | Is periodic interval and the mesaurement interval is both are same, setting the range from 10 second to 24 hours
SW6TS002 | The user can manualy trigger a measurement | Must have | Login as expert password:bernd | 1.Enter the URL 2..Login using Credintials 3.Click on Service 4.Select Measurement settings 5.Click on "Trigger Measurement" 6.Logout |  |  |  | After triggering measurement the page remain same, didn't not find any changes in the page
Testcase ID | Functional Requirement -SW8.Global Calibrations | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW8TS001 | A View is present that list the available global clibration and corrently the selected one | Must have | Login as expert password:bernd | 1.Enter the URL 2.Provide login credentials 3.Click on Service 4.Select Measurement Settings |  |  |  | Not found in the link
SW8TS002 | For every global calibration a short description can be viewed | Must have | Login as expert password:bernd |  |  |  |  | Not found in the link
SW8TS003 | It is possible to select the used Global Calibration and the parameters used from it | Must have | Login as expert password:bernd |  |  |  |  | Not found in the link
SW8TS004 | It is possible to upload new Global Calibration (via a file or via Internet) and delete existing ones | Must have | Login as expert password:bernd |  |  |  |  | Not found in the link
Testcase ID | Functional Requirement -SW9.Local Calibrations | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW9TS001 | The user must be able to calibrate either offset or offset and gradient for each parameter | Must have | Login as expert password:bernd | 1.Enter the URL 2.Provide Login Credentials 3.Click on Service 4.Select Measurement Settings 5.Click on enter sevice mode 6.Click on edit symbol beside parameter name 7.Navigates to parameter properties box 8.Check use of local calibration 9.Click on OK 10.Logout
SW9TS002 | The calibration can be performed by direct input/editing of the calibration coefficients | Must have | Login as expert password:bernd | 1.Enter the URL 2.Provide Login Credentials 3.Click on Service 4.Select Measurement Settings 5.Click on enter sevice mode 6.Click on edit symbol beside parameter name 7.Navigates to parameter properties box 8.Check use of local calibration 9.Click on OK 10.Logout
SW9TS003 | Old calibrations are kept and shown | Must have | Login as expert password:bernd | 1.Enter the URL 1.Login using Credentials 2.Click on Service 3.Select Measurement Settings 4.Click on enter sevice mode 5.Click on edit symbol beside parameter name 6.Navigates to parameter properties box 8.Check use of local calibration 7.Click on OK 8.Logout
SW9TS004 | User can select to use an old calibration | Must have | Login as expert password:bernd
Testcase ID | Functional Requirement -SW11.Referencing | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW10TS001 | The user must be enabled to create new energy refernce for distiled water and air | Must have | Login as expert password:bernd | 1.Enter the UR:L 1.Login using Credentials 2.Click on service 3.Click on Spectral Refrence 4.Create new Spectral reference 5.Provide the name and select reference type 6.Click on Ok 7.Logout | User should able to create a new energy reference for distilled water and air
SW10TS002 | A Comment shall be entered to describe the new refrence | Must have | Login as expert password:bernd | 1.Enter the URL 1.Login using Credentials 2.Click on service 3.Click on Spectral Refrence 4.Create new Spectral reference 5.Provide the comments 6.Click on Ok 7.Logout | User should able to provide comments
Testcase ID | Functional Requirement -SW12.Reference Selection | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW12TS001 | Old references are kept and the production reference is always kept | Must have | Login as expert password:bernd
SW11TS002 | A view is provided that shows the available references together with their comment and date of creation | Must have | Login as expert password:bernd
SW11TS003 | The user shall be able to a select the reference that is used by the device | Must have | Login as expert
Testcase ID | Functional Requirement -SW13.Configuration interface | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW13001 | It shall be possible to change network settings Ip settings (IPv4 and IPv6),host name | Must have | Login as expert password:bernd | 1.Enter the URL & valid credentials 2.Click on service tab 3.Click on device properties & click on edit settings 4.Under IP settings select the IP Mode as static 5.Enter the IP address (IPv4, IPv6) 6.Logut | IP setting can able to change in network settings
SW13002 | It shall be possible to change network settings VPN settings | Must Have | Login as expert password:bernd | Didn't get it need to check in the url
SW13003 | It shall be possible to change network settings NTP and time zone settings, including manual time settings | Must have | Login as expert password:bernd | 1.Enter the URL & valid credentials 2.Click on service tab 3.Click on device properties & click on edit settings 4.Under time setting section 5.Enter NTP & time zone settings 6.Logout | User can able to change the NTP & time zone settings
SW13004 | It shall be possible to change network settings as follows Email settings under SMTP settings | Must have | Login as expert password:bernd |  |  |  |  | Didn't get it need to check in the url
SW13005 | It shall be possbile to change network settings as follows Bluetooth settings (PAN/BENP profile, visibility,IP settings,enable/disable,only v3.1) | Must have | Login as expert password:bernd | 1.Enter the URL & valid credentials 2.
SW13006 | Configure settings for automatic data file upload (FTP/SFTP) 1.Server can be configured with IP, username, Password 2.Files may be transferred after every measurement or on time schedule | Must have | Login as expert password:bernd |  |  |  |  | Didn't get it & need to check
SW15007 | Settings on locale and user language | Must have | Login as expert password:bernd |  |  |  |  | Didn't get it & need to check
SW13008 | Sleep mode settings | Must have | Login as expert password:bernd |  |  |  |  | Didn't get it & need to check
SW13009 | Other configuration option referred by other requirements of this document | Must have | Login as expert password:bernd |  |  |  |  | Didn't get it & need to check
Testcase ID | Functional Requirement -SW25.RS485/GPIO settings | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW25001 | User can configure the legacy interface to act either as RS485 (half-duplex)or as two GPIO pins. | Must have | Login as expert password:bernd | 1.Enter the URL & valid credentials 2.Click on service tab & device properties 3.Under device settings section 4.Select GPIO mode 5.Logout |  |  |  | Didn't get it & need to check
SW25002 | In the case device is configured in the RS485 mode, the user can set : 1. Baudrate , parity and stop bot for the RS485 interface 2. Address and timeout for the Modbus communication | Must have |  | Didn't get it & need to check |  |  |  | Didn't get it & need to check
Testcase ID | Functional Requirement -SW15.Authentication | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW17001 | A login page is provided that accepts a user password. Also, a logout is available. | Must have | Login as expert password:bernd | 1.Enter the url & login page is opened 2.Enter the valid username & password 3.Click on perform login 4.Navigated to other pages 5.Logout | User should be logged in successfully
SW17002 | This password can be changed on the settings page | Must have | Login as expert password:bernd |  |  |  |  | Didn't find it in the link
SW17003 | At first time connection after production the user is asked to change the password. Also terms and conditions from s::can must be accepted at first usage | Must have | Login as expert password:bernd |  |  |  |  | Didn't find it in the link
SW17004 | After a period of inactivity the user is logged out. Probe returns to normal operation if it is stuck in some configuration process, e.g. calibration mode. The default time for inactivity is set to 30 minutes but can be changed by the user | Must have | Login as expert password:bernd |  |  |  |  | Didn't find it in the link
Testcase ID | Functional Requirement -Access Levels | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW16001 | The following access levels are supported 1.viewer 2.user 3.expert 4.s-can | Must have | Login as viewer password: Login as user password:bernd Login as expert password:bernd Login as s-can password: | 1.Enter the URL 2.Provide username as : viewer 3.Provide password as: 4.Logout as "viewer" 5.Provide username as :user 6.provide password:bernd 7.Logout as "user" 8.provide user name:expert 9.password:bernd 10.Logout as "expert" 11.provide username as:s-can 12.password: 13.Logout as s-can | Access level by following different roles 1.Verify the access level by login as "viewer". 2.Verify the access level by login as "user" 3.Verify the access level by login as "expert" 4.Verify the access level by login as "s-can"
SW16002 | The access levels have influence on which data can be shown and modified | Must have | Login as expert password:bernd
Testcase ID | Functional Requirement - User levels: views and permissions | Priority-1 | Pre-condition | Action & Brief Description | Expected result | Actual Result | Status | Comments
SW17001 | Access restrictions are as follow:
SW17002
Author Name
Reviewed By
Approved By